Supplementary Materials

EMI Platform: A Modular Embedded Measurement Architecture for Industrial SCADA Systems

These supplementary materials provide reproducibility information, experimental data provenance, frozen analytical results, and the materials required to reproduce the experimental validation of the EMI Platform.

# S1. Purpose and Scope

This document complements the main manuscript by documenting the experimental and analytical materials used for validation of the EMI Platform. It supports reproducibility while preserving the distinction between the generic EMI Platform architecture and the application-specific smart pyranometer implementation.

# S2. Reproducibility Repository

The source code, experimental datasets, analysis scripts, and reproducibility materials are released in the project repository: GitHub repository popase/EMI_Platform. The repository contains the EMI Platform implementation, firmware and gateway components, experimental series, and the frozen calibration and validation datasets used in the manuscript.

The released analytical data comprise 1,044 observations for calibration and 583 observations for independent validation. The repository also contains the analysis pipeline used to generate the reported results.

# S3. Experimental Data Provenance

The experimental data processing chain is: reference-instrument display images → OCR extraction → raw SPM72 readings → manual correction of OCR errors → timestamp synchronization with EMI Node measurements → derived analytical dataset.

Raw and corrected SPM72 readings are retained together with the source images in the reproducibility repository. Timestamp synchronization uses nearest-neighbor matching with a maximum allowed temporal difference of ±15 s.

# S4. Experimental Series Allocation

Series S01–S07 form the calibration dataset. Series S08-A, S08-B, and S09–S11 form the validation dataset. Series S08-C was excluded from the analytical datasets because it contains an identified experimental intervention/anomaly.

# S5. Frozen Analytical Datasets

The calibration and validation datasets were frozen before final model evaluation. The calibration dataset contains 1,044 observations, while the independent validation dataset contains 583 observations acquired on a separate measurement day. No validation observation was used for model estimation or model selection.

# S6. Frozen Calibration Model

The selected calibration model is linear and was estimated by ordinary least squares using exclusively the calibration dataset:

G = 6.550 × I − 37.577

Exact frozen coefficients: slope = 6.550183385298; intercept = −37.577182152696. Model identifier: EMI-LIN-2026-08-26.

A quadratic alternative was evaluated during calibration. Its improvement in calibration RMSE was approximately 0.42%, considered marginal relative to the additional parameter; therefore, the linear model was retained.

# S7. Validation Performance

| Metric | Validation result |
| --- | --- |
| N | 583 |
| MAE | 25.67 W/m² |
| RMSE | 36.07 W/m² |
| MBE | 4.44 W/m² |
| R² | 0.9791 |
| Pearson r | 0.9904 |

The validation results were obtained using the frozen linear model without recalibration on validation data. A total of 131 of the 583 validation observations lie outside the calibration irradiance interval; this is reported as an extrapolation-related limitation.

# S8. Communication Endurance

The industrial communication endurance test generated 10,586 Modbus RTU requests over approximately three hours. Of these, 10,585 requests were successful and one request resulted in a timeout. No observed CRC errors were recorded. The resulting successful transaction rate was 99.99%.

# S9. Reproduction Procedure

1. Deploy the released EMI Platform software and firmware components on the specified experimental hardware.

2. Acquire the reference SPM72 readings and EMI Node measurements using the documented acquisition workflow.

3. Apply the documented OCR correction and timestamp synchronization procedure.

4. Use the frozen calibration dataset and frozen linear model coefficients provided in the repository.

5. Process the independent validation dataset without modifying the model coefficients.

6. Recalculate the validation metrics and compare them with S7.

# S10. Reproducibility Boundaries

The supplementary materials support reproduction of the reported implementation, data processing, calibration, validation, and communication endurance results. They do not imply experimental validation of generic portability, scalability, or reuse across other microcontrollers, sensors, or industrial installations. These properties are architectural objectives of the EMI Platform.

Remote Configuration was not implemented or experimentally validated in the present study and is outside the validated scope of this article.

# External Supplementary Annexes

The following external annexes accompany the supplementary package and are referenced explicitly in the manuscript:

Annex A — Electrical and Wiring Schematic

Annex B — EMI Platform Interconnection Diagram

Annex C — Experimental FSM and Application Workflow

Annex D — Reproducibility Reference

Annex A documents the electrical and wiring implementation of the experimental EMI Node. Annex B documents the interconnection of the EMI Platform components. Annex C documents the experimental finite-state machine and application workflow. Annex D provides the detailed reproducibility reference for the released implementation, datasets, analysis pipeline, and reproduction procedure.