Supplementary Materials

EMI Platform: Reproducibility, Experimental Data, and Validation Materials

Supporting material for the associated manuscript

# S1. Purpose and scope

These Supplementary Materials provide reproducibility information and supporting analytical records associated with the experimental evaluation of the EMI Platform. They complement the manuscript without duplicating its architectural description or extending its scientific claims.

The EMI Platform is the primary subject of the study. The smart pyranometer is the application-specific experimental instantiation used to evaluate the platform.

# S2. Reproducibility repository

The complete reproducibility repository is maintained in the GitHub repository popase/EMI_Platform. It contains the analysis pipeline, firmware, experimental series, frozen calibration and validation datasets, and supporting documentation.

analysis/ — reproducible data-processing and diagnostic pipeline.

firmware/ — experimental EMI Node firmware.

gateway/ — gateway/FSM software.

series/ — the 13 final experimental measurement series.

analysis/EMI_Dataset/data/calibration_dataset.csv — frozen calibration dataset (N = 1044).

analysis/EMI_Dataset/data/validation_dataset.csv — frozen independent validation dataset (N = 583).

# S3. Experimental data provenance

Raw JPEG / OCR acquisition → spm72_readings_raw.csv → manual/visual OCR correction → spm72_readings.csv → synchronization with emi_measurements.csv → synchronized_measurements.csv → derived analysis dataset.

The original OCR output is preserved separately from corrected OCR values. The synchronized analytical data use the corrected OCR readings. Source measurement files are not modified by the analysis pipeline.

# S4. Experimental series and analytical allocation

Seven valid analytical sessions from 20–22 August 2026 were assigned to calibration. Five valid sessions or temporal segments from 24 August 2026 were reserved for independent validation. Other series were excluded from quantitative model evaluation according to the predefined validity decisions.

| Unit | Role | Status |
| --- | --- | --- |
| S01 | Calibration | VALID |
| S02 | Calibration | VALID – dynamic |
| S03 | Calibration | VALID |
| S04 | Calibration | VALID |
| S05 | Calibration | VALID |
| S06 | Calibration | VALID |
| S07 | Calibration | VALID – dynamic |
| S08-A | Independent validation | VALID |
| S08-B | Independent validation | VALID – dynamic |
| S08-C | Excluded | Not used quantitatively |
| S09 | Independent validation | VALID – dynamic |
| S10 | Independent validation | VALID |
| S11 | Independent validation | VALID – low irradiance |

# S5. Frozen calibration and validation datasets

| Dataset | Observations | Role | Status |
| --- | --- | --- | --- |
| Calibration | 1044 | Model estimation | FROZEN |
| Validation | 583 | Independent evaluation | FROZEN |

Validation observations were not used for model estimation or model selection. The split is performed by measurement day, with 24 August reserved for independent validation.

# S6. Application-specific measurement model

The final application-specific model is a linear relationship between PV current and reference irradiance:

G = 6.550 × I − 37.577

G is estimated irradiance in W/m² and I is PV current in mA. The coefficients were estimated exclusively from the frozen calibration dataset using ordinary least squares and frozen before independent validation.

Model identifier: EMI-LIN-2026-08-26.

The linear model was retained after comparison with a quadratic candidate: the quadratic model improved calibration RMSE by only approximately 0.42% while adding an additional parameter.

# S7. Model performance

| Metric | Calibration | Independent validation |
| --- | --- | --- |
| N | 1044 | 583 |
| MAE [W/m²] | 20.2539 | 25.6668 |
| RMSE [W/m²] | 50.7335 | 36.0706 |
| MBE [W/m²] | ≈0 | 4.4446 |
| R² | 0.9168 | 0.9791 |
| Pearson r | 0.9575 | 0.9904 |

# S8. Communication endurance evidence

| Measure | Result |
| --- | --- |
| Modbus requests | 10,586 |
| Successful transactions | 10,585 |
| Timeouts | 1 |
| Observed CRC errors | 0 |
| Success rate | 99.99% |

This test provides platform-level evidence for the stability of the embedded-to-gateway communication path under the tested conditions.

# S9. Reproduction procedure

From analysis/EMI_Dataset, after installing the dependencies listed in requirements.txt:

python 06_check_input.py

python 01_prepare_dataset.py

python 02_scatter_spm72_pvrear.py

python 03_timeseries.py

python 04_descriptive_report.py

python 07_series_diagnostics.py

Alternatively, 05_make_all.py runs scripts 01–04 sequentially, followed by 07_series_diagnostics.py.

# S10. Reproducibility boundaries

The generic analysis pipeline does not automatically refit the frozen article model.

The calibration/validation split is a methodological research decision and is not regenerated automatically.

Original experimental CSV files are preserved and are not modified by the analysis scripts.

The frozen datasets are treated as research artifacts and should not be silently regenerated or replaced.

The reported model and validation results are application-specific evidence supporting evaluation of the EMI Platform.

# Annex A. Repository structure relevant to reproducibility

README.md

analysis/EMI_Dataset/config.json

analysis/EMI_Dataset/common.py

analysis/EMI_Dataset/requirements.txt

analysis/EMI_Dataset/01_prepare_dataset.py … 07_series_diagnostics.py

analysis/EMI_Dataset/data/<experimental series>/

analysis/EMI_Dataset/data/calibration_dataset.csv

analysis/EMI_Dataset/data/validation_dataset.csv

firmware/EMI_Node/EMI_Node.ino

gateway/FSM/spm72_control.py

gateway/FSM/spm72_ocr.py

gateway/FSM/sync_series.py

# Annex B. Frozen analytical reference

Calibration: S01–S07, N = 1044.

Independent validation: S08-A, S08-B, S09, S10, S11, N = 583.

Model: linear OLS, estimated on calibration only.

Frozen coefficients: slope = 6.550183385298; intercept = −37.577182152696.

Validation: MAE = 25.67 W/m²; RMSE = 36.07 W/m²; MBE = 4.44 W/m²; R² = 0.9791; Pearson r = 0.9904.

Communication endurance: 10,585 successful transactions / 10,586 requests; one timeout; no observed CRC errors.

# Annex C. Scope boundary

Remote device configuration and lifecycle management are not implemented or experimentally validated in the present work. They therefore do not form part of the reproducibility evidence for the current article.

End of Supplementary Materials